Flame Photometry Protocol

Preparation of 1% nitric acid

- Work in the fume hood.

- Wear lab coat, eye protection, gloves and lower screen in front of face.

- Always add acid to water slowly, never water to acid.

- For 250 samples (250*5=1.25L), 1.5L of 1% nitric acid should be sufficient.

- If stock is 70% nitric acid, pour 69ml of MilliQ-water for every 1ml of nitric acid. (x 21.4 for 1.5L).

- If stock is 69% nitric acid, pour 68ml of MilliQ-water for every 1ml of nitric acid. (x 21.8 for 1.5L)

- Pour 21.8ml nitric acid to 1483.4ml MilliQ-water.

Preparation of samples:

- Put all samples in 10ml falcon tubes and determine fresh/dry weight of plant material.

- Add 3ml of 1% nitric acid to 10ml tubes of samples and close lids tightly.

- Put in oven at 70oC overnight.

- After 3 hours, shake gently without breaking the rosette to cover all in nitric acid.

- Remove samples and let them cool down on the bench before diluting.

Preparation of dilutions:

- Prepare dilutions in fresh MilliQ-water and make sure to be extremely precise in all measurements.

- Use 2 salt and 2 control eppendorfs to test the dilutions and standards before doing the rest of the samples.

- Flip the falcon tube to remove condensation, before pipetting.

- Prepare replicates for each sample, just in case.

- Make sure to avoid carrying over any plant-fibre to the dilutions.

Suggested dilutions: 	Salt 1:20 (1.90ml H2O + 100μL sample)	 Control 1:5   (1.6ml H2O + 400μL sample)

Suggested standards: 	Salt 500μM NaCl, 100μM KCl	             Control 10-100μM NaCl, 1000μM KCl

Operation of flame photometer:

- Fill two beakers with fresh MilliQ-water; the running solution and blank solution.

- Put the running solution under the U-tube and add water with droplet to the connecting tube.

- Check that the U-tube doesn’t have any air bubbles; wiggle it around if necessary.

- Check waste container and dispose of solution down the sink.

- Open “air” on the main valve and check that liquid is being drawn into the machine and that waste liquid is flowing out into waste bucket.

- Check at the back of instrument that pressure is approx. 11 psi.

- Open gas bottle by turning the valve all the way, then a half-turn back.

- Switch flame photometer and check that the flame has been lit, and then close the hatch.

- Turn computer on to record the measurements manually.

- Leave the running solution under the U-tube and let it run for 30-60 min before use.

Calibration:

- Shake the standard solution bottle before use, then add around 10ml to a beaker.

- Set system to dual and peak only.

- Put in blank solution, press “Blank”, wait till “100” is flashing, then remove blank solution.

- Put in selected standard solution and press “Cal” for Na and K and when “100” flashes, press “Cal” again.

- Press “Measure” and wait till it blinks to measure Blank (should be 0±2) and Standard (should be 100±2).

- Start measuring eppendorf samples and wipe the tube with tissue when changing solutions.

- Note: Correct measurements for Salt (20 - 120) and Control (10 - 20)

- Run standard as test every 12-24 Samples

- When switching between standard solutions, press “Blank” for 3 seconds to restart calibration.

Shut down and clean-up:

- Make sure the running solution is still under the U-tube, so the flame photometer doesn’t dry out.

- Turn gas valve off and wait until the error E61 is displayed and flame is out, then switch off and close air.

- Dispose of solution and plant material down sink with excess water.